

***Device Under Test***

|  |  |
| --- | --- |
| ***Asset Number:*** <AssetNumber> | ***Serial Number:*** <SerialNumber> |
| ***Location:*** <Location>  ***Manufacturer:*** <Manufacturer> | ***Model:*** <Model> |

***Test Equipment:***

|  |
| --- |
|  |

***Performance Test Summary***

|  |
| --- |
| **<result69>** |

***Signatures***

|  |  |
| --- | --- |
| Name: <Name> | Date: <Date> |

***Performance Test Results***

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Volume Accuracy | **Volume (mL) (± 5%)** | **Percentage (± 5%)** | | | **Results** | |
| Volume Test - 5 mL demand @ 300 mL/hr | **<volume1>** | **<percentage1>** | | | **<result1>** | |
| Calibration Factor for 5, 25, 125, 500 L/hr Final | **<calibrationfactor>** | | | | **<result2>** | |
| Pressure Calibration Data | | | **Results** | | | |
| Pressure Delta | | | **<result3>** | | | |
| Pressure Cap Reset to 255 | | | **<result4>** | | | |
| Mechanical Inspection | | | | **Results** | | |
| Physical Inspection | | | | **<result5>** | | |
| Battery Date Stamp - Replace Battery after two years | | | | **<result6>** | | |
| Time and Date Set | | | | **<result7>** | | |
| Initialisation and Power Up - No Error Messages | | | | **<result8>** | | |
| Display Window Intact and Readable | | | | **<result9>** | | |
| On/Alarm LED Green and Red Backlight, beeps green red | | | | **<result10>** | | |
| Keypad Follow Screen Prompt | | | | **<result11>** | | |
| Display Test - Red, Green. Blue, White | | | | **<result12>** | | |
| Sound Alarm Test | | | | **<result13>** | | |
| Door Sensor Test - Open or Close as Appropriate | | | | **<result15>** | | |
| Door Hinge and Latch Operates Smoothly | | | | **<result15>** | | |
| Air Sensor Test - Set Out Reads "AIR", Set In Reads "LIQUID" | | | | **<result16>** | | |
| Motor and Encoder Test - Reads "PASS" | | | | **<result17>** | | |
| Pressure Sensor Test - Set Out Reads 46-54 | | | | **<result18>** | | |
| Charger Test | | | **Serial Number** | | | **Results** |
| Serial Number CV | | | **<serial>** | | | **<result19>** |

***Comments:***

<Comments>