

***Device Under Test***

|  |  |
| --- | --- |
| ***Client:*** <Client> | ***Type:*** <Type> |
| ***Asset Number:*** <AssetNumber> | ***Serial Number:*** <SerialNumber> |
| ***Location:*** <Location>  ***Manufacturer:*** <Manufacturer> | ***Model:*** <Model> |

***Test Equipment:***

|  |
| --- |
|  |

***Performance Test Summary***

|  |
| --- |
| **<PerformanceTest>** |

***Signatures***

|  |  |
| --- | --- |
| Name: <Name> | Date: <Date> |

***Performance Test Results***

***Performance Verification:***

|  |  |
| --- | --- |
| Visual check | **<result1>** |

***Inspiratory Pressure Range Check:***

|  |  |
| --- | --- |
| Set range at 5 LPM (approx. 2-73cm H2O mbar) | **<result2>** |
| Set range at 8 LPM (approx. 4-73cm H2O mbar) | **<result3>** |
| Set range at 10 LPM (approx. 5-75cm H2O mbar) | **<result4>** |
| Set range at 15 LPM (approx. 10-77cm H2O mbar) | **<result5>** |

***Input Gas Flow Range:***

|  |  |
| --- | --- |
| Minimum at 5 LPM | **<result6>** |
| Maximum at 15 LPM | **<result7>** |

***PEEP Range Check:***

|  |  |
| --- | --- |
| Set range at 5 LPM (approx. 1-5cm H2O mbar) | **<result8>** |
| Set range at 8 LPM (approx. 1-9cm H2O mbar) | **<result9>** |
| Set range at 10 LPM (approx. 2-15cm H2O mbar) | **<result10>** |
| Set range at 15 LPM (approx. 4-25cm H2O mbar) | **<result11>** |

***Operating Time Check:***

|  |  |
| --- | --- |
| 5 LPM (approx. 80 minutes) | **<result12>** |
| 10 LPM (approx. 40 minutes) | **<result13>** |
| 15 LPM (approx. 26 minutes) | **<result14>** |

***Comments:***

<Comments>