

***Queensland Ambulance Service Report***

|  |  |
| --- | --- |
| ***Date:*** 13/04/2017 | ***Station:*** SPRINGFIELD |
| ***Location/Room:*** GARAGE  ***Registration Number:*** 915WBI | ***Vehicle Number:*** 6686 |

***Test Summary***

|  |  |
| --- | --- |
| ***List of Serviced Equipment:***  1x - Oxygen Reticulation: ( NVB0008695)  1x - Aspirator: (NVB0008694)  1x - Electric Suction: (NVB0008696)  2x - Regulator: (NVB0008691, NVB0008692)  1x - Multi-Flow Regulator: (NVB0008693) | ***List of Parts Replaced:***  3x - BODOK SEAL  1x - OVERHAUL KIT |

***Signatures***



|  |  |
| --- | --- |
| Name: Rommel Lapuz | Date: 13/04/2017 |

***Test Results***

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Multi-Flow Regulator | | |
| ***Asset Number****:* NVB0008693 | ***Serial Number:*** N/A | |
| ***Make****:* ALLIED | ***Model:*** MULTI-FLOW REGULATOR | |
| Pressure Test Static | | **630** |
| Check Flow Control & Full Flow | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:*  RECOMMEND OVERHAUL OR REPAIR | | |
| *Parts:*  BODOK SEAL | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Aspirator | | |
| ***Asset Number****:* NVB0008694 | ***Serial Number:*** 27431053 | |
| ***Make****:* OXYLITRE | ***Model:*** HIGH INJECTOR SUCTION UNIT | |
| Check Maximum Suction Function | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:* | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Regulator | | |
| ***Asset Number****:* NVB0008691 | ***Serial Number:*** 4710 | |
| ***Make****:* COMWELD | ***Model:*** O-SERIES REGULATOR | |
| Pressure Test Static | | **400** |
| Check Flow Control & Full Flow | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:*  BODOK SEAL | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Regulator | | |
| ***Asset Number****:* NVB0008692 | ***Serial Number:*** 3N13 | |
| ***Make****:* COMWELD | ***Model:*** O-SERIES REGULATOR | |
| Pressure Test Static | | **400** |
| Check Flow Control & Full Flow | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:*  BODOK SEAL  OVERHAUL KIT | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Oxygen Reticulation Failure Alarm | | |
| ***Asset Number****:* NVB0008695 | ***Serial Number:*** N/A | |
| ***Make****:* N/A | ***Model:*** N/A | |
| No Supply Left Ensure Alarm | | **Pass** |
| Alarm | | **Pass** |
| Alarm Clears | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:* | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Electric Suction | | |
| ***Asset Number****:* NVB0008696 | ***Serial Number:*** N/A | |
| ***Make****:* N/A | ***Model:*** N/A | |
| Check Condition & Smoothness of Valve | | **Pass** |
| Check Suction Max Negative KpA | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:* | | |