

***Queensland Ambulance Service Report***

|  |  |
| --- | --- |
| ***Date:*** 24/04/2017 | ***Station:*** BOONAH |
| ***Location/Room:*** NA  ***Registration Number:*** 074HDG | ***Vehicle Number:*** 6631 |

***Test Summary***

|  |  |
| --- | --- |
| ***List of Serviced Equipment:***  1x - Outlet Point: (NVB0008709)  1x - Aspirator: (NVB0008708)  1x - Electric Suction: (NVB0008711)  1x - Flowmeter: (NVB0008710)  2x - Regulator: (NVB0008706, NVB0008705)  1x - Multi-Flow Regulator: (NVB0008707) | ***List of Parts Replaced:***  3x - BODOK SEAL  2x - OVERHAUL KIT  1x - O2 NIPPLE |

***Signatures***



|  |  |
| --- | --- |
| Name: Rommel Lapuz | Date: 24/04/2017 |

***Test Results***

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Electric Suction | | |
| ***Asset Number****:* NVB0008711 | ***Serial Number:*** | |
| ***Make****:* | ***Model:*** | |
| Check Condition & Smoothness of Valve | | **Pass** |
| Check Suction Max Negative KpA | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:* | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Flowmeter | | |
| ***Asset Number****:* NVB0008710 | ***Serial Number:*** | |
| ***Make****:* | ***Model:*** | |
| Check Flow Control and Full Flow | | **Pass** |
| Check No Leaks | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:*  O2 NIPPLE | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Outlet Point | | |
| ***Asset Number****:* NVB0008709 | ***Serial Number:*** | |
| ***Make****:* | ***Model:*** | |
| Leak Test at Pipeline Pressure 400 KpA | | **Pass** |
| Visual Inspection | | **Pass** |
| Pressure Drop Test with Cyclinders Off | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:* | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Aspirator | | |
| ***Asset Number****:* NVB0008708 | ***Serial Number:*** | |
| ***Make****:* | ***Model:*** | |
| Check Maximum Suction Function | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:* | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Multi-Flow Regulator | | |
| ***Asset Number****:* NVB0008707 | ***Serial Number:*** 60738 | |
| ***Make****:* ALLIED | ***Model:*** MULTI-FLOW REGULATOR | |
| Pressure Test Static | | **500** |
| Check Flow Control & Full Flow | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:*  RECOMMEND OVERHAUL | | |
| *Parts:*  BODOK SEAL | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Regulator | | |
| ***Asset Number****:* NVB0008706 | ***Serial Number:*** 4710 | |
| ***Make****:* | ***Model:*** | |
| Pressure Test Static | | **400** |
| Check Flow Control & Full Flow | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:*  BODOK SEAL  OVERHAUL KIT | | |

|  |  |  |
| --- | --- | --- |
| ***Equipment Description****:*  Regulator | | |
| ***Asset Number****:* NVB0008705 | ***Serial Number:*** 3710 | |
| ***Make****:* | ***Model:*** | |
| Pressure Test Static | | **400** |
| Check Flow Control & Full Flow | | **Pass** |
| Leak Check | | **Pass** |
| Overall Test Result | | **Pass** |
| *Comments:* | | |
| *Parts:*  BODOK SEAL  OVERHAUL KIT | | |