


***Device Under Test***

|  |  |
| --- | --- |
| ***Client:*** adas  ***Asset Number:*** ad | ***Type:*** as213  ***Serial Number:*** as |
| ***Location:*** asd | ***Manufacturer:*** d |
| ***Model:*** das |  |

***Electrical Test Summary***

|  |  |  |  |
| --- | --- | --- | --- |
| **PASS** | | | |
| ***Test Details:*** | ***Test Instruments:*** |
| Date: 14/11/2017 | Ansur Version: ESA 620, v2.11 |
| Standard: AS NZS 3760 – Class 2 |  |

***Signatures***



|  |  |
| --- | --- |
| Name: Eric Chu | Date: 14/11/2017 |

***Electrical Safety Test Results***

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Mains Voltage – Live to Neutral | | | | |
| ***Results:***  Live to Neutral: | ***Value:***  239.3 VAC | ***High Limit:*** | ***Low Limit:*** | ***Standard:***  AS NZS 3760 – Class 2 |
| Mains Voltage – Neutral to Earth | | | | |
| ***Results:***  Neutral to Earth: | ***Value:***  1.6 VAC | ***High Limit:*** | ***Low Limit:*** | ***Standard:***  AS NZS 3760 – Class 2 |
| Mains Voltage – Live to Earth | | | | |
| ***Results:***  Live to Earth: | ***Value:***  239.3 VAC | ***High Limit:*** | ***Low Limit:*** | ***Standard:***  AS NZS 3760 – Class 2 |
| Protective Earth – Configuration Low | | | | |
| ***Results:***  Protective Earth Resistance: | ***Value:***  0.029 OHMS | ***High Limit:***  1 Ω | ***Low Limit:*** | ***Standard:***  AS NZS 3760 – Class 2 |
| Insulation Resistance - Configuration: Test Voltage: 500V | | | | |
| ***Results:***  Mains to Protective  Earth: | ***Value:***  >100 MOhms | ***High Limit:*** | ***Low Limit:***  1 MΩ | ***Standard:***  AS NZS 3760 – Class 2 |
| Earth Leakage Current – Normal Condition | | | | |
| ***Results:***  Normal Condition: | ***Value:***  30.7 uA | ***High Limit:***  1000 uA | ***Low Limit:*** | ***Standard:***  AS NZS 3760 – Class 2 |