

***Device Under Test***

|  |  |
| --- | --- |
| ***Asset Number:*** NVB0015513 | ***Serial Number:*** 14C353649 |
| ***Location:*** RFDS MT ISA  ***Manufacturer:*** BBRAUN | ***Model:*** BBRAUN CHARGER |

***Electrical Test Summary***

|  |  |  |  |
| --- | --- | --- | --- |
| **PASS** | | | |
| ***Test Details:*** | ***Test Instruments:*** |
| Date: 15/06/2017 | Ansur Version: ESA 620, v2.10 |
| Standard: AS NZS 3551 – Class 2 |  |

***Signatures***



|  |  |
| --- | --- |
| Name: Rommel Lapuz | Date: 15/06/2017 |

***Electrical Safety Test Results***

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Mains Voltage – Live to Neutral | | | | |
| ***Results:***  Live to Neutral: | ***Value:***  237.4 VAC | ***High Limit:*** | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Mains Voltage – Neutral to Earth | | | | |
| ***Results:***  Neutral to Earth: | ***Value:***  0.5 VAC | ***High Limit:*** | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Mains Voltage – Live to Earth | | | | |
| ***Results:***  Live to Earth: | ***Value:***  237.7 VAC | ***High Limit:*** | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Insulation Resistance - Configuration: Test Voltage: 500V | | | | |
| ***Results:***  Mains to Protective  Earth: | ***Value:***  >100 MOhms | ***High Limit:*** | ***Low Limit:***  1 | ***Standard:***  AS NZS 3551 – Class 2 |
| Touch Current– Normal Condition | | | | |
| ***Results:***  Normal Condition: | ***Value:***  0.1 uA | ***High Limit:***  100 | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Touch Current – Open Neutral | | | | |
| ***Results:***  Open Neutral: | ***Value:***  0.1 uA | ***High Limit:***  500 | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Touch Current – Open Earth | | | | |
| ***Results:***  Open Earth: | ***Value:***  0.1 uA | ***High Limit:***  500 | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Touch Current – Normal Condition, Reversed Mains | | | | |
| ***Results:***  Normal Condition,  Reversed Mains: | ***Value:***  0.1 uA | ***High Limit:***  100 | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Touch Current – Open Neutral, Reversed Mains | | | | |
| ***Results:***  Open Neutral, Reversed: | ***Value:***  0.1 uA | ***High Limit:***  500 | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |
| Touch Current – Open Earth, Reversed | | | | |
| ***Results:***  Open Earth, Reversed: | ***Value:***  0.1 uA | ***High Limit:***  500 | ***Low Limit:*** | ***Standard:***  AS NZS 3551 – Class 2 |