

***Device Under Test***

|  |  |
| --- | --- |
| ***Asset Number:*** 123 | ***Serial Number:*** 123 |
| ***Location:*** 123  ***Manufacturer:*** 12311 | ***Model:*** 123 |

***Test Equipment:***

***Signatures***



|  |  |
| --- | --- |
| Name: Rommel Lapuz | Date: 10/04/2017 |

***Performance Test Results***

***Performance Verification:***

|  |  |
| --- | --- |
| Pre OP Check | **Pass** |
| AED Function | **Pass** |
| Output @ Set 1-10 joules ±10% | **Pass** |
| Output @ Set 15 joules ±10% | **Pass** |
| Output @ Set 20 joules ±10% | **Pass** |
| Output @ Set 30 joules ±10% | **Pass** |
| Output @ Set 50 joules ±10% | **Pass** |
| Output @ Set 70 joules ±10% | **Pass** |
| Output @ Set 100 joules ±10% | **Pass** |
| Output @ Set 120 joules ±10% | **Pass** |
| Output @ Set 150 joules ±10% | **Pass** |
| Output @ Set 170 joules ±10% | **Pass** |
| Output @ Set 200 joules ±10% | **Pass** |
| Display Check | **Pass** |
| Sync Time < 60mSec | **Pass** |
| Post OP Check | **Pass** |
| ECG from Pads and Leads | **Pass** |
| Leads Off | **Pass** |
| Temp Check Set simulator to 40◦C. Displayed value 40 +/- 0.2◦C | **Pass** |
| Invasive P= 200mmHg. Displayed value 200 +/- 5 | **Pass** |
| NIBP Simulator Test (Calibration in Date) | **Pass** |
| SPO2 | **Pass** |
| CO2 Calibrated and Checked at 5% | **Pass** |

***Comments:***

asdasdsad