

***Device Under Test***

|  |  |
| --- | --- |
| ***Asset Number:*** as | ***Serial Number:*** das |
| ***Location:*** asd  ***Manufacturer:*** asd | ***Model:*** d |

***Test Equipment:***

|  |
| --- |
| HUATO HE704 THERMOMETER (SN HE20103717) |
| OXITEST PLUS 7 PULSE OXIMETER TESTER (SN D0S04090718) |

***Performance Test Summary***

|  |
| --- |
| **Pass** |

***Signatures***



|  |  |
| --- | --- |
| Name: Rommel Lapuz | Date: 4/07/2017 |

***Performance Test Results***

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Volume Accuracy | **Volume (mL) (± 5%)** | **Percentage (± 5%)** | | | **Results** | |
| Volume Test - 5 mL demand @ 300 mL/hr | **4.9** | **-2.00** | | | **Pass** | |
| Calibration Factor for 5, 25, 125, 500 L/hr Final | **140** | | | | **Pass** | |
| Pressure Calibration Data | | | **Results** | | | |
| Pressure Delta | | | **Pass** | | | |
| Pressure Cap Reset to 255 | | | **Pass** | | | |
| Mechanical Inspection | | | | **Results** | | |
| Physical Inspection | | | | **Pass** | | |
| Battery Date Stamp - Replace Battery after two years | | | | **Pass** | | |
| Time and Date Set | | | | **Pass** | | |
| Initialisation and Power Up - No Error Messages | | | | **Pass** | | |
| Display Window Intact and Readable | | | | **Pass** | | |
| On/Alarm LED Green and Red Backlight, beeps green red | | | | **Pass** | | |
| Keypad Follow Screen Prompt | | | | **Pass** | | |
| Display Test - Red, Green. Blue, White | | | | **Pass** | | |
| Sound Alarm Test | | | | **Pass** | | |
| Door Sensor Test - Open or Close as Appropriate | | | | **Pass** | | |
| Door Hinge and Latch Operates Smoothly | | | | **Pass** | | |
| Air Sensor Test - Set Out Reads "AIR", Set In Reads "LIQUID" | | | | **Pass** | | |
| Motor and Encoder Test - Reads "PASS" | | | | **Pass** | | |
| Pressure Sensor Test - Set Out Reads 46-54 | | | | **Pass** | | |
| Charger Test | | | **Serial Number** | | | **Results** |
| Serial Number CV | | | **123213asdasd** | | | **Pass** |

***Comments:***